## Slide 1
A Mind Controlled and User-Centered intelligent wheelchair
Mechatronics Engineering program CHS
2024 / 2025

## Slide 2
  Adham Ebrahem Abdelmonem
  Omar Mohamed Farouq Azab
  Adham Abd-Elfatah mohamed
  Zeyad Ahmed Abdelmonsef
Prof.Dr/Mohamed Ali ELsayed Agwa
Dr/Osama Talaat Elghonimy
A Mind Controlled and User-Centered intelligent wheelchair
2024 / 2025
Team Members:
  Yara Ahmed Maher
  Ali Yasser Ali El-Nady
  Hassan Yasser Hassan
  Salma Nabil Mohamed
  Maria Fady Fayek Nagib
  Ammar Yasser Abdelhafez
  Ahmed Ayman Ibrahem
Under Supervision of:
Mechatronics Engineering program CHS

## Slide 3
Programing
Table of Contents
Introduction
Mechanical
Electrical
Future Work
Testing
Conclusion

## Slide 4
Abstract
This project presents a practical and innovative solution in assistive technology to improve mobility for individuals with severe physical disabilities. It features a brain-controlled wheelchair that interprets brain signals to control direction.
A simple, real-time algorithm processes brainwave data without using machine learning, making the system easy to use and highly responsive. The developed system successfully controls wheelchair movement and includes ADAS support. This work sets the stage for future enhancements.

## Slide 5
Introduction
Throughout the design phase, a wide range of brain-based control options were explored but The decision-making process involved careful consideration of signal stability, user comfort, and real-time feasibility such as :
1-Alpha, Beta, Delta and Theta rhythms .
2- Muscular responses eye movements and jaw clenching .

## Slide 6
1-Process Motivation:
Their needs shaped every decision Simplicity, confidence, and ease of use were our top priorities ,So:
Brain-Computer Interface (BCI) technologies have emerged as a promising solution, allowing users to control devices using only brain activity. it becomes possible to translate neural activity into real-world commands.
This intelligent wheelchair places the user's needs at the center of its functionality. The system is designed to:
Provide hands-free control of wheelchair direction .

## Slide 7
7
Key advantage of the system is that it does not require training or user-specific calibration, unlike machine learning-based approaches which depend heavily on personalized data. It increases accessibility and makes the solution more practical .
2-Accessibility Without Personalization:

## Slide 8
3-Proof of Concept and Practicality:
This project was developed as a fully functional and integrated system, not just a prototype or conceptual model it took from previous projects and went further with inclusion of additional safety layers.
Milo
USC
Us

## Slide 9
  Overview of the mechanical design of the smart wheelchair.
  Key role in enhancing comfort and performance.
  Challenges in achieving balance and smooth movement.
  Transition from simulation to real-world application.
MECHANICAL DESIGN
9

## Slide 10
  Enhances comfort and performance.
  Quick responsiveness to brainwave signals.
  Integration with other systems.
Importance of the Mechanical
MECHANICAL DESIGN
10

## Slide 11
  Challenges in movement and balance.
  Quick system responsiveness.
  Integration with other systems.
11
MECHANICAL CHALLENGES

## Slide 12
  "Here, we will show a video demonstrating how we used SolidWorks to simulate the wheelchair’s movement, followed by its application in the real-world prototype."
12
SIMULATION TO REAL-WORLD APPLICATION

## Slide 13
13
SIMULATION TO REAL-WORLD APPLICATION

## Slide 14
14
MECHANICAL COMPONENTS AND DESIGN OVERVIEW

## Slide 15
  Key design considerations for strength and stability.
  Focus on weight distribution and balance.
  Use of lightweight but strong materials.
15
STRUCTURAL DESIGN

## Slide 16
  Type and size of wheels used.
  Role in maneuverability and surface compatibility.
  Design decisions to ensure stability and comfort.
16
WHEELS DESIGN AND FUNCTIONALITY

## Slide 17
  Type of motors used.
  Importance in movement and responsiveness.
  Integration with control system for brainwave interaction.
17
MOTORS AND ACTUATORS

## Slide 18
  Leather with a large foam layer for comfort.
  Pressure relief and better support.
  Ergonomic design for comfort and support.
18
Seat Comfort and Design

## Slide 19
  Custom acrylic box for controller and motor driver.
  Positioned above the front wheel and below the seat.
  Transparent for easy access and maintenance.
19
Control Box

## Slide 20
Properties | value
Yield strength: | 2.48168e+08 N/m^2
Tensile strength: | 4.82549e+08N/m^2
Elastic modulus: | 2e+11N/m^2
Poisson's ratio: | 0.32
Mass density: | 7,800kg/m^3
Shear modulus: | 7.6e+10N/m^2
Therma expansion coefficient: | 1.2e-05/Kelvin
Cast carbon steel was selected for the wheelchair frame due to its exceptional properties
20
Material
properties

## Slide 21
21
Four fixed-geometry fixtures were used in the study: two supporting the front castor wheels and two for the rear wheel axles.
STRUCTURE
ANALYSIS
After selecting cast carbon steel, a structural analysis was conducted to evaluate the design's ability to withstand loads and meet safety standards.
Fixtures

## Slide 22
22
Loads
A simulated load of 100 kg was applied, distributing approximately 980 N to the seat and 490 N to the backrest.
This allows us to evaluate how the chair performs under realistic conditions

## Slide 23
  Boundary conditions applied; results plotted for analysis.
The maximum observed stress was 4.993 × 10⁷ N/m², which is well within the yield strength limit, confirming the design's structural safety.
Stress
23
STUDY
RESULTS

## Slide 24
24
Displacement
Displacements below 0.485 mm are considered within the safe limit.

## Slide 25
25
Strain
A maximum strain of 1.526 × 10⁻⁴ is permissible for safe operation.

## Slide 26
26
The design is considered safe as long as the factor of safety exceeds the minimum value of 1.764.
Factor Of Safety

## Slide 27
  Actuators
  power
  Controllers
  Sensors
Neurosky EEG Sensor
Gyroscope sensor
Time of flight sensor
  EasyEDA simulation
THE MAIN COMPONENTS OF ELECTRICAL SYSTEM
27

## Slide 28
  Function: Generate the power to rotate the wheelchair in various directions.
  Specifications :
Quantity: 2 motors
Power: 110 Watts per motor
Voltage: 24 Volts
1- Actuators (DC Brushed Motors)
28
COMPONENTS

## Slide 29
1- Actuators (DC Brushed Motors)
Advantages :
  Readily available
  Cost-effective
  Simple construction
  Controller is easy and affordable
29
COMPONENTS

## Slide 30
2- POWER SUPPLY
- Batteries:
  Type: Gel/AGM lead-acid batteries
  Voltage: 12 Volts per battery
  Capacity: 30 Ampere-hour per battery
  Connection: Series connection (to provide 24 Volts)
30
COMPONENTS

## Slide 31
2- POWER SUPPLY
- POWER CHARGER :
  Current : 8 Amperes
  Voltage : 24 volt
  Function : Safely and efficiently charges the batteries
31
COMPONENTS

## Slide 32
3- CONTROLLERS
MOTOR DRIVE
MICROCONTROLLERS
32
COMPONENTS

## Slide 33
3- CONTROLLERS
  MICROCONTROLLERS
A microcontroller is a compact, integrated circuit designed to govern a specific operation in an embedded system.
33
COMPONENTS

## Slide 34
34
  MCU: STM32F446RE
  Core: ARM Cortex-M4 (32-bit) with Floating Point Unit (FPU)
  Clock Speed: Up to 180 MHz
  Flash Memory: 512 KB
  SRAM: 128 KB
.
STM32 NUCLEO F446RE FEATURES :
  MICROCONTROLLERS
COMPONENTS

## Slide 35
An affordable and flexible solution for testing concepts and building prototypes, offering a balance of performance and power efficiency.
Role in the wheelchair:
  EEG data processing from Neurosky.
  Motor control for the wheelchair.
  Sensors for obstacle detection and ADAS
35
  STM32 NUCLEO F446RE
MICROCONTROLLERS
COMPONENTS

## Slide 36
OVERVIEW:
  MOTOR DRIVE ( BTS7960 )
36
COMPONENTS
  The BTS7960 43A motor driver module is a powerful and reliable choice for controlling high-current DC motors in various embedded and automation systems.

## Slide 37
  MOTOR DRIVE ( BTS7960 )
FEATURES :
37
COMPONENTS
  Supports up to 43A continuous current
  Operating voltage: 5.5V to 27V
  Full speed & direction control via PWM
  Dual H-Bridge for efficient control
  Built-in protections:
  Over-temperature
  Over-current
  Short-circuit

## Slide 38
  MOTOR DRIVE ( BTS7960 )
38
COMPONENTS
Role in the wheelchair:
  Enables safe and smooth motion of the wheelchair.
  Provides reliable, high-power control.
  Essential for mobility and precision in electric wheelchair systems

## Slide 39
Neurosky EEG Sensor
  Blink Detection : Strength and frequency of blinks trigger forward and backward movement.
  Brain-Computer Interface (BCI) : Detects brain activity non-invasively for natural control.
39
SENSORS
EEG Sensor for Forward and Backward Motion

## Slide 40
Sensor Detecting signal
40

## Slide 41
  Head Tilt Left : Wheelchair moves left when head tilts beyond set angle.
  Head Tilt Right : Wheelchair moves right with corresponding tilt detection.
  MEMS 3-aixs accelerometer and 3-axis gyroscope values combined
GYROSCOPE SENSOR:
TYPE USED : MPU6050
41
HEAD MOVEMENT FOR DIRECTIONAL CONTROL

## Slide 42
42
Accelerometer Only
Gyroscope Only
Both Together (IMU)
Can detect tilt and orientation but is noisy during movement
Combines both to give stable, precise orientation and motion tracking
Detects rotation only, cannot detect tilt or gravity
HEAD MOVEMENT FOR DIRECTIONAL CONTROL

## Slide 43
  Four ToF Sensors : Placed at all corners for 360-degree obstacle detection.
  measures the distance to an object by emitting infrared light
  Automatic Safety Response : Sensors trigger automatic stop when obstacles are detected.
43
ADVANCED OBSTACLE DETECTION
TYPE USED : TOF400C-VL53L1X

## Slide 44
44
Advantages of ToF Sensors vs. Ultrasonic Sensors
Precision
ToF sensors provide more accurate distance measurements.
Faster Response
as it is Light-based sensing
They offer quicker obstacle detection for safer navigation.
Smaller Size
Compact design fits easily on all wheelchair corners.
Robustness
Less affected by environmental noise compared to ultrasonic sensors

## Slide 45
  Gyroscope sensor
45
EasyEDA simulation
  EEG Sensor
  two motors
  two motor drivers
  four time of flight sensor
  microcontroller STM32
components:

## Slide 46
PROGRAMMING AND the
Brain-Controlled Interface
46
  Signal Processing and Feature Extraction
  Wheel Chair Control System
Overview

## Slide 47
Signal Processing and Feature Extraction
• Noise filtering and signal enhancement.
47
  There are several types of Noise filtering:
- Notch Filtering
- smoothing
- Band-PassFiltering

## Slide 48
48
  There are several types of Noise filtering.
- Notch Filtering
Signal Processing and Feature Extraction

## Slide 49
49
  There are several types of Noise filtering.
- SMOOTHING
Signal Processing and Feature Extraction

## Slide 50
  There are several types of Noise filtering.
- Band-PassFiltering
50
Signal Processing and Feature Extraction

## Slide 51
  Extracting key features like EEG patterns or specific brainwave frequencies.
51
Signal Processing and Feature Extraction

## Slide 52
•This diagram illustrates the process of acquiring, processing, and classifying EEG signals to generate commands for controlling the brain-controlled wheelchair.
52
Wheel Chair Control System

## Slide 53
The control system uses brainwave data (blink & attention) or head tilt detected by a gyroscope to control a wheelchair.
  Blinks trigger forward, backward, and stop commands
  Attention levels control left, right, and stop actions
  Head tilt via gyroscope allows intuitive left/right navigation as a second mode
  Integrated with an ADAS (Advanced Driver Assistance System) to enhance safety and autonomy.
Wheel Chair Control System
53

## Slide 54
Blink strength is mapped to forward, backward counts, and specific blink patterns trigger movement commands like 'F' (Forward), 'B' (Backward), or 'S' (Stop).
54
Wheel Chair Control System
  EEG Control – Blink Commands

## Slide 55
Depending on attention level, the system sends 'R' (Right), 'L' (Left), or 'S' (Stop); low attention stops the system entirely.
55
Wheel Chair Control System
  Brainwave Control – Attention Levels

## Slide 56
56
Wheel Chair Control System
Gyroscope-Based Head Tilt Control
  Uses gyroscope to detect head tilt direction
  Intuitive directional control:
• Tilt left: Turn left
• Tilt right: Turn right
  Alternative to brainwave control; supports user flexibility

## Slide 57
57
Wheel Chair Control System
  Advanced Driver Assistance System (ADAS)
  Integrated Time-of-Flight (ToF) sensor for distance measurement
  Continuously monitors space in front of the wheelchair
  Safe Distance Threshold: > 50 cm
  If distance ≤ 50 cm, the wheelchair automatically stops
  Ensures real-time collision prevention in crowded or dynamic areas
ADAS – Obstacle Avoidance with ToF Technology

## Slide 58
58
Wheel Chair Control System
  Safety Features Overview
  Time-of-Flight Sensor (ADAS):
• Continuously monitors for obstacles
• Automatically stops if object is ≤ 50 cm away
  Controlled Motor Commands via STM32:
• Prevents sudden or unintended movements
• Only responds to valid, recognized signals
  Dual Control Modes:
• Allows fallback if one mode fails (Brainwave ↔ Gyroscope)
  Hands-Free Operation:
• Reduces user fatigue and risk of control errors
  Comprehensive Safety Measures in Wheelchair Control System

## Slide 59
The electrical and control system is the backbone of the project, responsible for interpreting neural signals and converting them into precise motor actions.
59
CONTROL PROCESS

## Slide 60
MIND TO MOTION: SMART EEG WHEELCHAIR WITH REAL-TIME SAFETY
60

## Slide 61
Step 1 - Signal Acquisition
Objective: Capture real-time brain activity.
  Device: MindWave Mobile 2 EEG sensor.
  Signal Types: Attention, Meditation, and Blink strength.
  Connection: Bluetooth to Python-based PC/server.
MIND TO MOTION: SMART EEG WHEELCHAIR WITH REAL-TIME SAFETY
61

## Slide 62
Step 2 - Signal Processing
Convert raw EEG signals into usable control commands.
  Language: Python.
  Method: Parsing and filtering data using json format.
  Filtering: Data cleaning and validation (blink strength, attention thresholds).
  Output: Clean values forwarded for decision-making.
MIND TO MOTION: SMART EEG WHEELCHAIR WITH REAL-TIME SAFETY
62

## Slide 63
MIND TO MOTION: SMART EEG WHEELCHAIR WITH REAL-TIME SAFETY
63

## Slide 64
Step 3 - Decision Making
Objective: Translate brain signal values into movement commands.
  Device: STM32 Microcontroller.
  Logic:
    Blink → Move Forward.
    High Blink → Stop or Move Backward.
    Gyroscope → Turn Left/Right.
MIND TO MOTION: SMART EEG WHEELCHAIR WITH REAL-TIME SAFETY
64

## Slide 65
Objective: Physically move the wheelchair.
  Components: Dual Motor Drivers.
  Commands: Received via UART from STM32.
  Control: Smooth start, stop, and direction change based on brain signals.
Step 4 - Motor Actuation
MIND TO MOTION: SMART EEG WHEELCHAIR WITH REAL-TIME SAFETY
66

## Slide 66
MIND TO MOTION: SMART EEG WHEELCHAIR WITH REAL-TIME SAFETY
67
Objective: Maintain safe and stable operation.
  Sensors: 4 TOF Distance Sensors and 1 MPU6050 Gyroscope.
  Functions:
    Obstacle Detection (front/back/side).
    Fall and Tilt Detection via Gyroscope.
  Response: Auto-stop or rerouting if unsafe conditions are detected.
Step 5 - Safety and Feedback

## Slide 67
MIND TO MOTION: SMART EEG WHEELCHAIR WITH REAL-TIME SAFETY
68

## Slide 68
69
Testing
from mind to motion

## Slide 69
CONCLUSION

## Slide 70
  The project demonstrates a significant leap forward in assistive technology, offering people with severe motor disabilities a real, hands-free mobility solution
  With successful testing of EEG signal translation, obstacle avoidance, and user comfort, the wheelchair is both functional and scalable.
CONCLUSION
73

## Slide 71
74

## Slide 72
  Healthcare Sensors Integration: We aim to incorporate health-monitoring sensors such as heart rate monitors, body temperature sensors, and pulse oximeters. These will help in detecting fatigue, stress, or medical emergencies while using the chair.
75
FUTURE WORK

## Slide 73
FUTURE WORK
  GPS Navigation System: The future model will feature GPS modules for real-time location tracking and auto-routing, enabling autonomous indoor and outdoor navigation.
76

## Slide 74
FUTURE WORK
  Mobile App Connectivity:
A dedicated app will be developed to allow remote monitoring, caregiver alerts, and control features through smartphones.
77

## Slide 75
FUTURE WORK
  Artificial Intelligence (AI) Integration: Machine learning models will be trained on user behavior to adapt the chair’s response, personalize performance, and reduce false triggers.
78

## Slide 76
REFERENCE
1. Niedermeyer, E., & da Silva, F. L. (2004). Electroencephalography: Basic
Principles, Clinical Applications, and Related Fields (5th ed.). Lippincott
Williams & Wilkins.
https://www.wolterskluwer.com/en/solutions/ovid/niedermeyers?electroencephalography-basic-principles-clinical-applications-and-related?fields-8611
2. Lebedev, M. A., & Nicolelis, M. A. L. (2006). Brain–machine interfaces: Past,
present and future. Trends in Neurosciences, 29(9), 536-546.
https://pubmed.ncbi.nlm.nih.gov/16859758/
3. He, H., Wu, D., & Li, Y. (2016). A brain-controlled wheelchair system for
disabled people. IEEE Transactions on Biomedical Engineering, 63(10), 2302–
2311.
https://www.researchgate.net/publication/224137450_A_Brain_Controlled_Wh
eelchair_to_Navigate_in_Familiar_Environments
79

## Slide 77
4. Zhu, X., Zhang, S., & Zhang, D. (2015). A brain-controlled wheelchair using
EEG signals. IEEE Access, 3, 1696-1704.
https://pmc.ncbi.nlm.nih.gov/articles/PMC8473300/
5. Liu, Y., Sourina, O., & Tan, J. (2008). Real-time EEG neurofeedback for brain–
computer interface. Proceedings of the International Conference on Computer
Graphics and Digital Imaging, 1-4. https://journalofethics.ama?assn.org/article/ethical-and-social-challenges-brain-computer-interfaces/2007-
6. Zhang, Y., & Wang, J. (2014). A novel EEG-based brain-computer interface
design for wheelchair control. Journal of Neural Engineering, 11(4), 046002.
https://www.worldscientific.com/doi/abs/10.1142/S0129065714500142?srsltid
=AfmBOorGV4X20sAvhYxQWiollM5wKrsbAhp9WyZUXwIE7h96dszepvrv
80
REFERENCE

## Slide 78
81
fOR YOUR ATTENTION
Thank You